## Slide 1: Lempeä johdatus sähköiseen impedanssitomografiaan Gentle introduction to Electrical Impedance Tomography
Samuli Siltanen 2025

## Slide 2: In EIT, one injects electric currents, measures voltages, and makes an image of the inside
The EIT image resonctruction task is a nonlinear and ill-posed inverse problem.
This presentation aims to demonstrate these facts using a simple pixel-based diffusion model.
There is a connection to partial differential equations through a finite difference approximation.

## Slide 3: EIT:ssä syötetään sähkövirtoja, mitataan jännitteitä ja muodostetaan kuva sisuksista
EIT:n kuvanmuodostustehtävä on epälineaarinen ja virheherkkä
käänteisongelma.
Esityksessä havainnollistetaan EIT-mittausta
diffuusiomallin avulla.
Yhteys osittaisdifferentiaaliyhtälöihin tulee differenssimenetelmän kautta.

## Slide 5: Voltage-to-current measurement, PDE-style

## Slide 6: Current-to-voltage measurement, PDE-style

## Slide 7: Termejä suomeksi
Capacitance = kapasitanssi Conductivity = johtavuus Conductance = konduktanssi Current = virta Ground = maa Grounded = maadoitettu Impedance = impedanssi Inductance = induktanssi Resistivity = sähkönvastus Resistor = vastus Tomografia = viipalekuvaus* Voltage = jännite
*vaikka EIT:n tapauksessa ei kuvata viipaleita

## Slide 8: First observation: EIT does not make sense in dimension one

## Slide 9: The starting point of EIT is Ohm’s Law: U=ZI
Here U is electric voltage, I is electric current and Z is impedance. Impedance resists the movement of electricity.
For example, in a basic RCL circuit with alternating current of moderate frequency we would have a complex-valued impedance
where R is resistance, L is inductance, C is capacitance and w is the angular frequency of a time-harmonic electric current. We can measure the impedance by driving a current of known amplitude through the circuit and measuring the voltage needed to do that.

## Slide 10: We assume zero frequency and consider only voltage, current, and resistance
R
This is a resistor with resistance R, measured in Ohm (W). The bigger is R, the harder it is for electric current to pass through the resistor.

## Slide 11: We apply an electric voltage U, measured in volts (V)
U>0
U=0
-
+
R
Ground

## Slide 12: There will be an electric current I, measured in amperes (A), going from + to -
R
U>0
U=0
I
By convention, electrons flow against the current, from – to +
-
+

## Slide 13: The crucial equation is R = U/I
R
U>0
U=0
I
-
+

## Slide 14: In EIT studies it is more convenient to work with conductivity instead of resistance
Let’s denote the conductance by G = 1/R. The unit of C is Siemens, with S = 1/W.
Recall the crucial equation R = U/I. A calculation gives I=GU.
Conductivity s(x) is defined pointwise, and its unit is S/m.

## Slide 15: Modelling variable conductivity with grids
Think of many square pixels connected to each other, with constant conductance inside every pixel. Then we can approximate a continuously changing s(x) by a piecewise constant model.

## Slide 16: We consider a simple one-dimensional electric problem
Voltages v_i are at the electrodes, accessible from the surface.
Voltages u_j are voltages inside the grid cells making up the rod of conductive material. Conductivity is now one everywhere.

## Slide 17: We feed 1 mA of current into the rod and take it out at the opposite end
As a result, we get voltages at the electrodes and also inside the rod. In our simple computational model, voltage is constant inside a cell.

## Slide 18: We feed 1 mA of current into the rod and take it out at the opposite end
The right electrode is grounded, so voltage there is zero.

## Slide 19: The voltages in numbers
We measure 4 volts at the left electrode.

## Slide 20: Now let’s have a non-uniform conductivity distribution
Electric current now can flow more easily between cells 1 and 2, as the conductivity is 5 instead of 1.
1
5
1
1

## Slide 21: Now the cells 1 and 2 have almost the same color
It is so easy for current to flow between cells 1 and 2 that the voltages even out almost completely.

## Slide 22: Here is the previous result as numbers: cells 1 and 2 have almost the same voltage
The right electrode is always zero, so our voltage measurement is on the right. Here we get 3.2, which is less than the 4 we got earlier.

## Slide 23: In electrical impedance tomography we would only observe at the boundary
Feeding 1 mA of current through the rod results in measuring voltage
3.2 V.
3.2

## Slide 24: What if we change the location of the anomalous conductivity?
Electric current now can flow more easily between cells 2 and 3, as the conductivity is 5 instead of 1.
1
1
5
1

## Slide 25: Now the cells 2 and 3 have almost the same color
It is so easy for current to flow between cells 2 and 3 that the voltages even out almost completely.

## Slide 26: Here is the previous result as numbers: cells 2 and 3 have almost the same voltage
We again measure 3.2.

## Slide 27: Electrical impedance tomography does not see any difference between the two options
3.2
5
?
?
Our boundary measurements have just two numbers: current amplitude and measured voltage on the right.

## Slide 28: We arrive at the conclusion that one-dimensional EIT is impossible
Applying current with I amperes results in measuring U volts. Then we can calculate the total resistance as
R = U/I.
All we know is the total conductance (integral of s(x)):
G = 1/R = I/U.

## Slide 29: How about EIT in dimension two?

## Slide 30: Current-to-voltage measurement, PDE-style

## Slide 31: Two-dimensional grid of pixels, with constant conductivity 1
There are two kinds of variables:
Voltages “v_i” measured at the 8 electrodes,
Voltages “u_j” at the 9 internal pixels. We cannot access them in reality; here in the model we are able to observe them.

## Slide 32: Do we simply get the same situation than in the one-dimensional case?
0
0
0
0
0
0
0
0
0
0
0
0
0
0
Spoiler: NO. This is not how physics works.

## Slide 33: To answer the previous question, let’s study the rules of electric current flow
u
uW
uE
uN
uS
Involved are voltages in a center pixel and its four neighbors

## Slide 34: Add conductivities between pixels.
u
uW
uE
uN
uS
Each s tells how easily current can flow through the edge.
sN
sS
sW
sE

## Slide 35: Consider the flow of electric current from left to the center
u
uW
uE
uN
uS
The voltage difference driving the current is uw – u. If uw is larger than u,
then current flows into the center.
How large is the current?
IW = sW(uw – u).

## Slide 36: Conservation of charge requires that the four currents sum up to zero
u
uW
uE
uN
uS
We get the equation
0 = IW+ IS+ IE+ IN
= sW(uW – u)
+ sS(uS – u)
+ sE(uE – u)
+ sN(uN – u)

## Slide 37: Let’s see if these numbers add up
Recall that our inspiration was the 1D case

## Slide 38: We need to solve a matrix equation for u&v
=

## Slide 39: Correct solution of the 2D problem

## Slide 40: Next point: EIT is nonlinear
1
0.5
1.5
First conductivity
Second conductivity
1.5
0.5
1

## Slide 41: Same current feeds for all
First conductivity
Second conductivity
Conductivity equals 2 everywhere

## Slide 42: Voltages as numbers. They do not add up!
+
=
+
=

## Slide 43: Next point: more resolution, and current patterns & voltage matrices

## Slide 44: Let’s study a 9x9 case. Left: constant conductivity 1, right: two inclusions
0.1
3
1

## Slide 45: We apply a set of skip-4 current patterns
Each column is a current pattern: a vertical 32-vector listing the currents through each of the electrodes.
Every pattern has one electrode injecting current into the domain. That current comes out from another electrode four steps away from the injecting one.

## Slide 46: Away from the inclusions, the two targets produce very similar data

## Slide 47: When the current passes through the resistive inclusion a lot, we see a big difference

## Slide 48: Difference is not as great with the conductive inclusion

## Slide 49: Two different conductivities yield two different voltage matrices with the same current matrix

## Slide 50: Information about the inclusions is encoded into the voltage matrices in a nonlinear way
Voltage matrix: conductivity 1 Voltage matrix: inclusions Difference
Here the three images have the same scale.

## Slide 51: Now we apply trigonometric current patterns
Each column is a current pattern: a vertical 32-vector listing the currents through each of the electrodes.

## Slide 52: The first current pattern probes the domain with quite evenly distributed power

## Slide 53: Current pattern 9

## Slide 54: Last, most oscillatory pattern 31

## Slide 55: Two different conductivities yield two different voltage matrices with the same current matrix

## Slide 56: Information about the inclusions is encoded into the voltage matrices in a nonlinear way
Voltage matrix: conductivity 1 Voltage matrix: inclusions Difference
Here the three images have the same scale.

## Slide 57: Next point: it is easier to probe near the boundary than to see deep inside

## Slide 58: Four conductivities, with the same inclusion but located differently

## Slide 59: We choose the inclusion to be more conductive than background
5

## Slide 60: Here are the voltage matrices resulting from trig currents

## Slide 61: Here are the voltage matrices resulting from trig currents
Voltage matrix of constant conductivity 1

## Slide 62: Subtracting the cond=1 voltage matrix from the measurements reveals effect of the inclusion
The deeper the inclusion is located, the less effect it has on EIT data. For the theoretical foundation of this fact, see
David Isaacson, "Distinguishability of Conductivities by Electric Current Computed Tomography," in IEEE Transactions on Medical Imaging, vol. 5, no. 2, pp. 91-95, June 1986, doi: 10.1109/TMI.1986.4307752.

## Slide 63: Further points
Does adding electrodes give arbitrarily accurate reconstructions? No, since the accuracy of each electrode should improve as well.
Optimal current patterns in the “distinguishability” sense of David Isaacson
2D versus 3D: what happens “off the plane”?
Complex impedance (alternating current)
Electrode models (gap, shunt, complete)
Boundary shape may be irregular or even unknown
How about reconstructions? What is the most basic method? Linearization, Calderón’s method, back-propagation (Barber-Brown)
Let’s collect the fundamental papers in one place
At some point: Wikipedia update

## Slide 65: The following stuff is draft material

## Slide 66
2
2
2
2
2
2
2
2
2

## Slide 67
2
2
2
2
2
2
2
2
2
2
2
2